BCM Testing Procedures

OMICRON

Before you start testing the BCM board, make sure to have the board wired correctly.

TESTING COMTRADE FILES

Hardware Setup

Analog Inputs:

In order to wire the analog inputs properly, please follow the next steps:

First, make sure that you have three of the cable shown in the image. It can be any color as long as it looks as the one in the picture.

Second, you should have three 75 amps MAGNETLABs that look as the one shown in the image. Afterwards, you must place each cable inside of the MAGNETLAB and make sure that they all phase a positive direction.

Third, when wiring you should do it as the one in the picture. First, connect the start of the cable to the current output one and the end of this to the current output N. Then, in TB6 as shown in the image, connect the MAGNETLAB as follows, the white cable in pin 1 and the black cable in pint 2. Second, connect the start of the cable to the current output 2 and the end of this to the current output N (right on top of the previous cable as shown in the image). Then, in TB6, connect the MAGNETLAB as follows, the white cable in pin 3 and the black cable in pin 4. Third, connect the start of the cable to the current output 3 and the end of this to the current output N (right on top of the previous cable). Then, in TB6, connect the MAGNETLAB as follows, the white cable in pin 5 and the black cable in pin 6.

Finally, the end result of the wiring should look as the following image.

Digital Inputs:

In order to wire the digital inputs properly, please follow the following steps:

First, make sure you have six cables as that shown in the image. I would suggest having three read and three black to be able to distinguish them easily.

Second, you will need at least three jumps that look as that shown in the image.

Third, when wiring you should do it in the following manner. First, connect one of the jumpers from the AUX DC red terminal, or plug, to the Binary Output 1(second row). Then, plug another jumper from the Binary Output 1(second row, right on top of the previous one) to the Binary Output 2(second row). Lastly, plug another jumper from the Binary Output 2(second row, right on top of the previous one) to the Binary Output 3 (second row). Afterwards, proceed to plug a cable, preferably red, from the Binary Output 1(first row) to the Digital Input 1 in TB1. Next, plug a cable, preferably black, from the AUX DC black terminal to the Digital Input 2 in TB1. Then, plug a cable, preferably red, from the Binary Output 2(first row) to the Digital Input 5 in TB1. Next, plug a cable, preferably black, from the AUX DC black terminal (right on top of the previous cable) to the Digital Input 6. Lastly, plug a cable, preferably red, from the Binary Output 3(first row) to the Digital Input 7 in TB1. Next, plug a cable, preferably black, from the AUX DC black terminal (right on top of the previous cable) to the Digital Input 8 in TB1.

Note -> Create tables

Software Setup

FileZilla: In order to be able to check for the BCM COMTRADE files in more efficient manner, I would suggest downloading the software FileZilla to access the BCM files in the following manner.

First, download the software FileZilla from the following online source, http://download.cnet.com/FileZilla/3000-2160_4-10308966.html, and install it in your machine. This software version is comparable with Windows XP/Vista/ and 7.

Second, proceed to enter the UNIX user account and password provided to you to access the BCM settings. You must do this as shown in the image below.

Lastly, once you have logged into the BCM account you can proceed to configure the software according the BCM specification and test.

BCM Configuration:  For the purpose of our testing the BCM must be configured as following.

First, go to the OMICROM Test Universe icon and double click on it. Once you have the window shown in the figure x, in “Text Module” click on “Advanced TransPlay” located on the forth row of the first column as highlighted on figure x.

If you already have the BCM Configuration file already in your in pc you must import it by clicking on import. Then, click on the file (BCMconfig…), click on “open”, and once this window has disappeared you should click on “apply” to update recent changes.

On the other hand, if you do not have the BCM Configuration file for this test you should create it in the following manner. On the “Hardware Configuration” window click on “Analog Outputs” and make sure that the settings look exactly as those in figure x. Then, click on “Binary / Analog Inputs” and make sure that it looks as figure y. Lastly, click on “Binary Output” and make sure that it looks as that in figure z.  After all these settings are done, you should proceed to export the configurations already made and save it as “BCMconfiguration+date”. In this way, you do not have to make the same configurations every time you need to test the BCM.

AuxDC Configuration: In order for you to test the BCM, you must first turn on the AuxDC to an appropriate voltage value. In this test we will be using only 110volt, and less you have been told otherwise you should steak to satisfy the BCM specifications. To turn on the AuxDC you must do the following. First, go back to the OMICROM Test Universe main window. Here, you need to below “Configuration Modules” and click on “AuxDC Configuration”. Once you have this window opened, proceed to change the voltage to 110 volts and click on “Apply Configuration” as shown in figure x.

Note: Before you proceed to the Test Procedure, you must make sure that the light, in OMICROM, of the AUXDC section is on. If this did not turn on after the above procedure you must close any window opened, from the OMICROM Test Universe application, and repeat the above procedure.

Test Procedure:

Before you start testing the BCM you must make sure that you have the BCM testing files. These files will provide you with the data necessary to test the BCM in different scenarios.

If you have not created a “TransPlay” file to test the BCM, you must do in the following manner.  Import the testing files from the directory in which you have saved these. An example is shown in figure x, in this case I have saved the testing files in “My Documents.” Afterwards, you must save the testing file as a Transplay file by going on to “File” and clicking on “Save As” as shown in figure y, and save the file with a name that represent the testing file as shown in figure z.

If you already have a “TransPlay” file created for the testing file, then proceed to test the unit.

To test the BCM you need to first open the TransPlay file, this files contains the three phases’ currents data and their behavior as data. Once you have already opened the file, you must click on “Start/Continue” button right on the header of the TransPlay window. You will hear a clicking sound from the OMICROM, this means that current of the phases A, B, and C have already been injected to the BCM. If you did not hear the clicking sound from the OMICROM, it could mean two things; first, the OMICROM is not able to communicate with the pc and therefore it cannot communicate with the board. Second, the OMICROM unit does not have the application “TransPlay” activated and therefore you can’t continue this test any further.

LABVIEW

SF6 TEST

Hardware Setup:

In order to test the SF6 application in the BCM board, you must have equipment such as a host to pressurize the tank, a bucket with enough water (about ¾), and tank as the one shown in figure x.

Now, you can proceed to setup the hardware to test the SF6 modules in the BCM board.

First, place a polyurethane tube from the PSI meter to the bottom of the bucket as shown in figure x. Make sure that about ¾ of the bucket is filled with water, this way you can check the speed of the leak from the tank.

Second, place a polyurethane tube from the front switch of the tank to the SF6 sensor. Also, make sure that the back of this switch is covered by place polyurethane tube with

Test Procedures:

IP Network Configuration

BCMname login: root

Password: x x x x

root@BCMname:~ ifconfig

eth0      Link encap:Ethernet  HWaddr 00:0C:C6:79:DE:5E

inet addr:10.10.92.20  Bcast:10.10.92.255  Mask:255.255.255.0

UP BROADCAST RUNNING MULTICAST  MTU:1500  Metric:1

RX packets:0 errors:0 dropped:0 overruns:0 frame:0

TX packets:0 errors:0 dropped:0 overruns:0 carrier:0

collisions:0 txqueuelen:1000

RX bytes:0 (0.0 B)  TX bytes:0 (0.0 B)

lo        Link encap:Local Loopback

inet addr:127.0.0.1  Mask:255.0.0.0

UP LOOPBACK RUNNING  MTU:16436  Metric:1

RX packets:16 errors:0 dropped:0 overruns:0 frame:0

TX packets:16 errors:0 dropped:0 overruns:0 carrier:0

collisions:0 txqueuelen:0

RX bytes:1272 (1.2 KiB)  TX bytes:1272 (1.2 KiB)

root@BCMname:~ ifconfig eth0 down

root@BCMname:~ ifconfig eth0 10.75.24.36

[  147.070000] eth0: Freescale FEC PHY driver [Generic PHY] (mii_bus:phy_addr=1:00, irq= -1)

root@BCMname:~ ifconfig eth0 hw ether 00:0C:C6:79:DE:89

root@BCMname:~ ifconfig eth0 up

root@BCMname:~ ifconfig

eth0      Link encap:Ethernet  HWaddr 00:0C:C6:79:DE:89

inet addr:10.75.24.36  Bcast:10.255.255.255  Mask:255.0.0.0

UP BROADCAST RUNNING MULTICAST  MTU:1500  Metric:1

RX packets:0 errors:0 dropped:0 overruns:0 frame:0

TX packets:0 errors:0 dropped:0 overruns:0 carrier:0

collisions:0 txqueuelen:1000

RX bytes:0 (0.0 B)  TX bytes:0 (0.0 B)

lo        Link encap:Local Loopback

inet addr:127.0.0.1  Mask:255.0.0.0

UP LOOPBACK RUNNING  MTU:16436  Metric:1

RX packets:16 errors:0 dropped:0 overruns:0 frame:0

TX packets:16 errors:0 dropped:0 overruns:0 carrier:0

collisions:0 txqueuelen:0

RX bytes:1272 (1.2 KiB)  TX bytes:1272 (1.2 KiB)

root@BCMname:~ cd /opt/Qualitrol/config/

root@BCMname:/opt/Qualitrol/config vi STATICIP

10.75.24.36

255.255.255.0

10.75.24.255

…

~

root@BCMname:/opt/Qualitrol/config vi MAC.txt

00:0C:C6:79:DE:89

…

~

root@BCMname:/opt/Qualitrol/config ifconfig

eth0      Link encap:Ethernet  HWaddr 00:0C:C6:79:DE:89

inet addr:10.75.24.36  Bcast:10.255.255.255  Mask:255.0.0.0

UP BROADCAST RUNNING MULTICAST  MTU:1500  Metric:1

RX packets:0 errors:0 dropped:0 overruns:0 frame:0

TX packets:0 errors:0 dropped:0 overruns:0 carrier:0

collisions:0 txqueuelen:1000

RX bytes:0 (0.0 B)  TX bytes:0 (0.0 B)

lo        Link encap:Local Loopback

inet addr:127.0.0.1  Mask:255.0.0.0

UP LOOPBACK RUNNING  MTU:16436  Metric:1

RX packets:16 errors:0 dropped:0 overruns:0 frame:0

TX packets:16 errors:0 dropped:0 overruns:0 carrier:0

collisions:0 txqueuelen:0

RX bytes:1272 (1.2 KiB)  TX bytes:1272 (1.2 KiB)

root@BCMname:/opt/Qualitrol/config bin

-sh: bin: not found

root@BCMname:/opt/Qualitrol/config cd ../BCMLite/bin/

root@BCMname:/opt/Qualitrol/BCMLite/bin ls

BCMLitexmls         configfs            mmwr

DNPSlavePoints.csv  data_logger         modbusdata

Guardian            dmserver            ntpd

SPIDataManager      dnpserver           ntpq

SPIDataManager.old  ipc_process         stamp

analysis            ipc_tags            xml_log_fs

bcm_time            ka.sh               xml_print

comtradefs          log_process

config_tool         mmrd

root@BCMname:/opt/Qualitrol/BCMLite/bin ./config_tool

Please select an instance for configuration

1: ipc

2: Guardian

3: time

4: xml_print

5: dm

6: spidatamanager

7: modbusdata

8: dnpsrv

9: analysis

10: data_logger

11: comtradefs

12: alarm_proc

13: wika_sf6

14: Exit

Select instance to configure: 2

Please select a configuration item

1: config/company_name

2: config/station_name

3: config/system_name

4: config/equipment_name

5: config/equipment_type

6: config/DHCP

7: config/ethaddr

8: config/ethmask

9: config/ethbcast

10: config/DNSserver[0]/domain

11: config/DNSserver[0]/server[0]

12: config/DNSserver[0]/server[1]

13: config/DNSserver[1]/domain

14: config/DNSserver[1]/server[0]

15: config/DNSserver[1]/server[1]

16: config/route[0]/destination

17: config/route[0]/gateway

18: config/route[0]/IPmask

19: config/route[1]/destination

20: config/route[1]/gateway

21: config/route[1]/IPmask

22: config/route[2]/destination

23: config/route[2]/gateway

24: config/route[2]/IPmask

25: config/route[3]/destination

26: config/route[3]/gateway

27: config/route[3]/IPmask

28: Exit

Select item to configure: 7

Current value: "10.10.92.20"

Please enter new value: 10.75.24.36

Please select a configuration item

1: config/company_name

2: config/station_name

3: config/system_name

4: config/equipment_name

5: config/equipment_type

6: config/DHCP

7: config/ethaddr

8: config/ethmask

9: config/ethbcast

10: config/DNSserver[0]/domain

11: config/DNSserver[0]/server[0]

12: config/DNSserver[0]/server[1]

13: config/DNSserver[1]/domain

14: config/DNSserver[1]/server[0]

15: config/DNSserver[1]/server[1]

16: config/route[0]/destination

17: config/route[0]/gateway

18: config/route[0]/IPmask

19: config/route[1]/destination

20: config/route[1]/gateway

21: config/route[1]/IPmask

22: config/route[2]/destination

23: config/route[2]/gateway

24: config/route[2]/IPmask

25: config/route[3]/destination

26: config/route[3]/gateway

27: config/route[3]/IPmask

28: Exit

Select item to configure: 8

Current value: "0.255.255.255.0"

Please enter new value: 255.255.255.0

Please select a configuration item

1: config/company_name

2: config/station_name

3: config/system_name

4: config/equipment_name

5: config/equipment_type

6: config/DHCP

7: config/ethaddr

8: config/ethmask

9: config/ethbcast

10: config/DNSserver[0]/domain

11: config/DNSserver[0]/server[0]

12: config/DNSserver[0]/server[1]

13: config/DNSserver[1]/domain

14: config/DNSserver[1]/server[0]

15: config/DNSserver[1]/server[1]

16: config/route[0]/destination

17: config/route[0]/gateway

18: config/route[0]/IPmask

19: config/route[1]/destination

20: config/route[1]/gateway

21: config/route[1]/IPmask

22: config/route[2]/destination

23: config/route[2]/gateway

24: config/route[2]/IPmask

25: config/route[3]/destination

26: config/route[3]/gateway

27: config/route[3]/IPmask

28: Exit

Select item to configure: 9

Current value: ""

Please enter new value: 10.75.24.255

Please select a configuration item

1: config/company_name

2: config/station_name

3: config/system_name

4: config/equipment_name

5: config/equipment_type

6: config/DHCP

7: config/ethaddr

8: config/ethmask

9: config/ethbcast

10: config/DNSserver[0]/domain

11: config/DNSserver[0]/server[0]

12: config/DNSserver[0]/server[1]

13: config/DNSserver[1]/domain

14: config/DNSserver[1]/server[0]

15: config/DNSserver[1]/server[1]

16: config/route[0]/destination

17: config/route[0]/gateway

18: config/route[0]/IPmask

19: config/route[1]/destination

20: config/route[1]/gateway

21: config/route[1]/IPmask

22: config/route[2]/destination

23: config/route[2]/gateway

24: config/route[2]/IPmask

25: config/route[3]/destination

26: config/route[3]/gateway

27: config/route[3]/IPmask

28: Exit

Select item to configure: 28

Configuration of Guardian:guard/config completed, closing update session

Please select an instance for configuration

1: ipc

2: Guardian

3: time

4: xml_print

5: dm

6: spidatamanager

7: modbusdata

8: dnpsrv

9: analysis

10: data_logger

11: comtradefs

12: alarm_proc

13: wika_sf6

14: Exit

Select instance to configure: 14

Configuration program terminated

root@BCMname:/opt/Qualitrol/BCMLite/bin neroute: SIOCDELRT: No such process

tstatroute: resolving 10.7510.2092.201

-rn

Kernel IP routing table

Destination     Gateway         Genmask         Flags   MSS Window  irtt Iface

10.75.24.0      0.0.0.0         255.255.255.0   U         0 0          0 eth0

root@BCMname:/opt/Qualitrol/BCMLite/bin add route default 10.75.24.1

-sh: add: not found

root@BCMname:/opt/Qualitrol/BCMLite/bin route add default 10.75.24.1

route: SIOCADDRT: No such device

root@BCMname:/opt/Qualitrol/BCMLite/bin

barebox 2011.03.0-dirty (Nov  5 2013 - 11:28:43)

Board: Karo TX28

Malloc space: 0x46800000 -> 0x47000000 (size  8 MB)

Stack space : 0x467f8000 -> 0x46800000 (size 32 kB)

running /env/bin/init...

booting kernel of type zimage from /dev/disk0.2

loaded zImage from /dev/disk0.2 with size 1887264

commandline: console=ttyAMA0,115200 tx28_base=tx28stk5 rootwait ro ip=none noinitrd root                                                                                                           fstype=ext2 root=/dev/mmcblk0p5

arch_number: 3043

Uncompressing Linux... done, booting the kernel.

…

done.

running rc.d services...

starting udev

mounting tmpfs at /dev

creating static nodes

starting udevd...done

waiting for devices...done

syslogd starting

klogd starting

hwclock: can't open '/dev/misc/rtc': No such file or directory

run-parts: /etc/rc.d/S16dropbear exited with code 1

starting sshd...done

starting network interfaces...

ip: SIOCSIFFLAGS: Cannot assign requested address

loading modules

OSELAS(R)-BCM-2013.11.05 (ptxdist-2011.04.0/2013-11-05T13:05:28-0500)